CERTIFICATE OF CALIBRATION
Antikernel Labs
PO Box 4665
10355 NE Valley Rd
Rollingbay, WA 98061-0665
https://www.antikernel.net/
sales@antikernel.net

This document certifies that the instrument described has been verified to comply with all published datasheet specifications using traceably calibrated equipment.

Instrument Description	2 GHz Transmission Line Probe
Model	AKL-PT1
Serial Number	0012

Calibration Performed By	A. Zonenberg	Date	2020-08-05 23:39
Test Conditions	21ºC, 42% R	H	Cal Due	2021-08-05
# Calibration Standards
Type	Manufacturer	Model	Serial	Cal due date
Multimeter	Rohde & Schwarz	HMC8012	36174847	2021-04-15
Vector Network Analyzer	Pico Technology	PicoVNA 106	09335	2021-04-17
SOLT Standard	Pico Technology	TA345	SF:296	2021-04-17
Oscilloscope	Teledyne LeCroy	WaveRunner 8404M-MS	LCRY4254N20447	2021-07-16
Pulse Generator*	Leo Bodnar	SMA Pulse Generator	P00718	2021-08-01

* Used as signal source only, not as measurement standard
# Test Results
Unless otherwise specified measurements are of probe body only, with cable and fixture de-embedded.
Test	Minimum	Measured	Maximum	Unit	Result
DC resistance	449.75	450.29	450.75	Ω	PASS
S21 (1 MHz)	-20.48	-20.46	-20.42	dB	PASS
S21 (500 MHz)	-20.85	-20.59	-20.35	dB	PASS
S21 (1.0 GHz)	-21.10	-20.70	-20.35	dB	PASS
S21 (1.5 GHz)	-21.45	-21.33	-20.75	dB	PASS
S21 (2.0 GHz)	-21.60	-22.10	-22.45	dB	PASS
Bandwidth (-23.5 dB)	2.25	2.54	2.60	GHz	PASS
Rise time (20-80%) w/ cable	118	120	129	ps	PASS

# Typical Waveform
# Nominal 40 ps risetime pulse including the provided cable if applicable
# S-Parameter Data
Machine readable S2P files for de-embedding may be downloaded at:
https://www.antikernel.net/downloads/AKL-PT1/caldata/0012/

## Insertion Loss (tip ground, across 50Ω termination)
tipground.s2p S21

## Insertion Loss (leaf ground, across 50Ω termination)
leafground.s2p S21

## Insertion Loss (Z-ground, across 50Ω termination)
zground.s2p S21

## Insertion Loss (flex ground, across 50Ω termination)
flexground.s2p S21

## Return Loss (tip ground, across open circuit)
zin.s2p S11